규  격  서 [A]

관세분류번호 H. S. NO	품목번호 ITEM NO.	품목 및 규격 Description	단위 Unit	수량 Quantity
		C-arm X선 촬영기	식	1

A. 특징 및 장점

1. Compact and Ergonomic Design for easy operation.

2. Superior Imaging Quality.

3. Dose Reduction through Various Pulsed Fluoroscopy Modes.

4. Wide SID(Source-to-Image Distances), Wiser operation.

5. Seamless compatibility with DICOM 3.0

B. 사양

1. C-arm Mobile Stand

1) Orbital rotation	: 135° (-45°∼+90°)

2) Vertical movement	: 500mm

3) Horizontal movement	: 200mm

4) Pivot rotation	: ± 225°

5) Panning movement	: ± 12.5°

6) Free space	: 780mm

7) SID(Source-to-Image Distances)	: 1000mm

8) C-arm Depth	: 720mm

2. X-ray Generator

1) X-ray tube	: Stationary anode

2) Focal spot zone	: 0.5 / 1.5mm

3) Anode heat capacity	: 32kJ (45kHU)

4) Maximum output	: 2.2kW

5) Fluoroscopy kV range	: 40 ~ 110kV

6) Fluoroscopy mA range	: Continuous fluoro 0.2 - 4mA

: Pulsed flouro 0.2 -  4mA

Boost : 5 - 10mA

7) Pulse mode	: 3 - 5 - 7.5 - 15 - 30 f/s

8) Radiography kV range	: 40 ~ 110kV

9) Radiography mA range	: 20mA

10) Radiography mAs range	: 0.4 ~ 100mAs

11) Collimator

- Automatically open and close by a motor

- Movement	: 24mm x 24mm

- Inherent filtration	: 2.5mmAL

3. High definition Image Intensifier (9")

1) Nominal entrance field size	: 9 Inch (23cm)

2) Field size	: 3 Field (23-15-10cm)

3) Resolution	: Over 48lp/cm in centor

4) Integrated laser positioning device on Image Intensifier

4. Digital Video Camera

1) Video format	: 1280(H) × 1024(V) Pixel, (VESA)

2) Image format	: 14bit B/

3) Output	: 1028×1024 pixel (30 frames/sec)

5. Monitor

1) Monitor	: 42“ Full HD LED monitor or more

2) Display	: Dual display

3) Resolution	: 1920 x 1080

6. Digital Image Processing

1) Digital Imaging Processing

- Multi-layered digital filter for optimized fluoroscopy and radiographic images

- Automatic brightness control not affected by moving subject

- Accurate feedback system supported by the Auto Exposure Control

2) Image data interface	: Ethernet (Gigabit)

3) DR single examination	: 1K × 1K, 14bits

4) Workstation

- Image processing software and the user interface ensuring easy usage and

management with Auto windowing, Digital Compensation Filtering and Fine

Auto Brightness Control

- OS	: Microsoft Windows 7, 32bit

- CPU	: Intel Core 2duo or more

- Hard disk	: 200 GB or more

- Memory	: DDR2 SDRAM 2G or more

- DVD R	: DVD-RW

5) Two control panels on the both sides of C-arm

7. DICOM 3.0 Interface

- Storage

- Worklist

8. Operating manual

C. 구성품

1. C-arm mobile stand	-1-

2. X-ray generator	-1-

3. High definition Image Intensifier 9"	-1-

4. Digital Video Camera	-1-

5. Monitor	-1-

6. Digital Imaging processing	-1-

7. DICOM 3.0 Interface	-1-

8. Operating manual	-1-

9. 납가운, 납장갑, 목보호대	-2-

10. 이동형 차폐막	-1-

D. 기타

1. Installation

- Supplier or contractor should be responsible for installation of the equipment.

2. Warranty

- Supplier or contractor should be responsible for free of charge after service whatever there is any defect or any trouble for three years after being accepted by customer.

규  격  서 [B]

관세분류번호 H. S. NO	품목번호 ITEM NO.	품목 및 규격 Description	단위 Unit	수량 Quantity
		C-arm X선 촬영기	식	1

A. 특징 및 장점

- I.I (Image Intensifier)  : 9 Inch

- Monitor Type         : 34" LCD

1. Generator

1) Type : High Frequency Inverter (80KHz)

2) Power Output : 5.3 KW, Active Air Cooling System

3) Fluoroscopy mode (투시모드)

- Continuous Mode    : 40 ~ 120 kV / 0.2 ~ 6.0 mA

- Pulse Mode         : 40 ~ 120 kV / 0.2 ~ 10.0 mA 1,2,4,8,15 fps

- Snap Shot Mode    : 10 ~ 50 mA

- ABC (자동밝기조정)  : 모든 모드에서 가능함. (Continue/Pulse)

- Anatomical Mode    : 4가지 모드 지원 (L-spine, C-spine, 관절부, 사지)

4) Radiographic mode (촬영모드)

- kV / mA / mAs      : 40 ~ 110kV / Max. 20mA / 1.0 ~ 100mAs

2. Tube & Mono Tank

- Anode Type         : Stationary

- Heat capacity       : 50.7 kHU(35.5kJ)

- Tank Heat capacity  : 1.5 MHU

- Focal Spot Size     : 0.5 mm / 1.8 mm

- Target Angle        : 16º

3. Collimator

-Type               : IRIS / 4 Blades Shutter

- Opening / Closing : 2 way

- Rotation           : 180º

4. Image Intensifier

- Input FOV size, inch (cm)  : 9"

- Nominal Entrance field size : 9 / 6 / 4.5 inch

- Central resolution (Lp/mm)  : 52 / 58 / 68

- Contrast ratio (10%)        : 30 : 1

- Laser pointer

5. Image Rotation

1) 방식           : 디지털 Rotation

2) 각도           : Limitless (Digital)

3) Rotation Guide : 모니터상에 Guide 표시

4) 반전           : 좌우, 상하

5) Memory Frame

- 기본저장 매수           : 70,000

- 저장방식                : 컴퓨터 HDD or USB Memory

- 동영상저장 (CINE Loop) : 초당 30 Frame, 30분 이상 기록가능.

- 동영상 플레이어         : 반복, 구간설정, 구간반복, 영상 캡쳐 가능.

6. Monitor

- Type                 : LCD

- Size (inch, Q'ty)      : 34 inch * 1대

- Resolution            : 1280 x 1024

- Brightness (cd/m2)   : 800

7. Dimensions

- SID (mm)             : 1,000

- Free Space (mm)     : 800

- Depth (mm)           : 720

- Orbital rotation (mm)  : 150º (90º/-60º)

- Horizontal travel (mm) : 200

- Vertical travel (mm)    : 500 (Motorized)

- Pivot roration          : ±125º

- Swiveling              : ±25º

8. 기타

- ZENIS   : 영상저장, DICOM 4.0, PACS(Q&R), DSA(옵션)

- 입력전원 : Single Phase, 220Vac, 5.3kW

B. 구성품

1. Generator	-1-

2. Tube & Mono tank	-1-

3. Collimator	-1-

4. Image intensifier	-1-

5. Image processing unit	-1-

6. Monitor	-1-

7. Main unit	-1-

8. Operating manual	-1-

9. 납가운, 납장갑, 목보호대	-2-

10. 이동형 차폐막	-1-

C. 기타

1. Installation

- Supplier or contractor should be responsible for installation of the equipment.

2. Warranty

- Supplier or contractor should be responsible for free of charge after service whatever there is any defect or any trouble for three years after being accepted by customer.